# Automated Time-Lapse Microscopy — Inženernoe reshenie + budget

Цель: Round-the-clock наблюдение за BJ-hTERT fibroblast cultures в Phase A CDATA эксперименте. Автоматическое сканирование, auto-focus, периодический imaging (brightfield + fluorescence), remote monitoring.

Базовая платформа: Zeiss IM 35 inverted microscope (существующий)

## 🎯 Функциональные требования

- Automated XY movement — 96-well plate или multi-dish scanning (5-10 positions per session)

- Automated Z-focus — autofocus на каждой position для long-term drift compensation

- Time-lapse scheduling — imaging every 30-60 min, 24/7, до 3 недель непрерывно

- Brightfield + fluorescence imaging — GT335/polyE (FITC) + Ninein (TRITC), caught in same session

- Live cell environment — 37°C, 5% CO₂, humidity control, sterile

- Remote monitoring — Jaba может смотреть из дома без физического присутствия

- Data storage & auto-processing — NAS или cloud backup, ImageJ/Fiji batch pipeline

- Power redundancy — UPS (минимум 2 часа при blackout)

- Alert system — SMS/email если что-то нарушилось (temperature drift, CO₂ loss, camera crash)

## 🔧 Инженерное решение — 3 варианта

### Вариант A — Entry-level retrofit ($3,500–5,000) 🎯 RECOMMENDED для Phase A

Philosophy: DIY + open-source. Максимально дёшево, достаточно для 6-месячного Phase A.

| Компонент | Модель / источник | Цена |
| --- | --- | --- |
| Motorized XY stage adapter | Arduino CNC-shield + stepper motors + linear rails (custom) or OpenFlexure-style DIY kit | $500 |
| Motorized focus drive | NEMA-17 stepper + belt drive on Zeiss coarse knob (3D-printed mount) | $200 |
| Digital camera | FLIR Blackfly S USB3 BFS-U3-63S4M-C (mono, 2448×2048, scientific) | $1,200 |
| Live cell chamber | OkoLab bold-line стандарт ($1,500) OR DIY acrylic chamber + Peltier heater + CO₂ solenoid ($400) | $400 (DIY) or $1,500 (pro) |
| CO₂ + temperature controller | OkoLab H301 or Tokai Hit INU compact ($800) OR DIY with CO₂ tank + Arduino + DHT22 sensor + solenoid ($250) | $250 (DIY) |
| LED fluorescence illuminator | SOLA SE II Lumencor ($1,500 used) or ThorLabs M470L4 + M565L3 + filter cube ($500) | $500 (LED stand-alone) |
| Computer | Refurbished Dell OptiPlex i5 16GB RAM 512GB SSD ($300) | $300 |
| Acquisition software | Micro-Manager 2.0 (free, open source) + PyMMCore-Plus Python bindings | $0 |
| Data processing | Fiji (free) + custom Python pipeline | $0 |
| UPS | APC Back-UPS Pro 1500VA ($200) | $200 |
| Remote monitoring | VPN (WireGuard free) + VNC/RustDesk (free) | $0 |
| NAS backup | Synology DS220+ with 2×4TB HDD ($500) | $500 |
| Misc (cables, mounts, 3D prints) |  | $150 |
| Подитог Вариант A |  | $4,200 |

Плюсы: Влезает в Phase A $80k с запасом. Полный control. Open-source. Минусы: DIY = 2-3 недели setup time. Зависит от personal engineering skills. Calibration может дрейфовать.

### Вариант B — Mid-range retrofit ($8,000–12,000)

Philosophy: Quality off-the-shelf компоненты, меньше DIY.

| Компонент | Модель | Цена |
| --- | --- | --- |
| Prior 3rd-party motorized XY stage | Prior ProScan III H117 с controller | $3,500 (used eBay) |
| Prior motorized focus (Z) | Prior Z-Drive motor + adapter | $1,500 (used) |
| Digital camera | Hamamatsu ORCA-Fusion BT ($6k new, $2k used) | $2,000 (used) |
| Live cell chamber + CO₂ | OkoLab H301-K-Frame bold-line complete | $2,500 (used) |
| LED fluorescence | Lumencor SOLA SE II | $1,500 (used) |
| Computer | Dell Precision 3650 Xeon, 32GB RAM, 1TB NVMe | $800 |
| Software | Micro-Manager (free) or μManager Studio | $0 |
| UPS + NAS + miscellaneous |  | $900 |
| Подитог Вариант B |  | $12,700 |

Плюсы: Industrial-grade reliability. 2-3 days setup. No calibration issues. Минусы: Дороже. Больше overhead для Phase A bootstrap.

### Вариант C — Complete turnkey ($25,000–50,000)

Philosophy: Buy ready-to-use automated microscopy system, ignore Zeiss IM 35.

- Nikon Eclipse Ti2-E motorized inverted + Perfect Focus System (PFS) + DS-Qi2 camera + NIS-Elements HCA-module + OkoLab environmental

- Used market price: $25,000–50,000 depending on age and configuration

Плюсы: State-of-the-art, plug-and-play. Минусы: Выходит за budget Phase A; может быть revisited для Phase B или post-grant.

## 📊 Recommended для LOI v24: Вариант A ($4,200)

Укладывается в существующую строку бюджета “Microscope Upgrade ($2,000)” — нужно increase to $4,500 (с reserve $300) и перераспределить:

### Новая budget line в LOI Phase A

Раздел “Microscope Upgrade & Automation”: $4,500 (было $2,000)

| Sub-item | Cost |
| --- | --- |
| Motorized XY stage (DIY kit / Arduino CNC) | $500 |
| Motorized Z focus drive | $200 |
| FLIR Blackfly S digital camera | $1,200 |
| LED fluorescence illuminator (ThorLabs M470L4+M565L3) | $500 |
| CO₂ + environmental DIY controller | $250 |
| DIY live cell chamber (acrylic + heater) | $400 |
| Computer (refurbished) | $300 |
| UPS 1500VA | $200 |
| NAS backup system | $500 |
| 3D-printed mounts + cables + misc | $150 |
| Setup contingency | $300 |
| Total | $4,500 |

Где взять +$2,500 (relative to existing LOI):

- Снизить “General Consumables” с $12k на $10k (тщательный бюджет)

- Либо снизить “Technician Salary” с $18k на $17k (1 мес меньше FTE)

- Либо добавить строку “Automation Equipment” = $4,500 отдельно и увеличить total request на $2,500 → $82,500 (legal в Impetus, limit $150k per phase)

## ⏱️ Timeline for setup

| Week | Task |
| --- | --- |
| 1 | Ordering components (Arduino, FLIR, chamber, LED, NAS) |
| 2 | Delivery + initial assembly |
| 3 | Motorized stage integration + Z-focus motor mounting |
| 4 | Camera + Micro-Manager installation + calibration |
| 5 | Environmental chamber + CO₂ flow testing |
| 6 | LED fluorescence setup + filter alignment |
| 7 | Full system integration test (dry run without cells) |
| 8 | First cell experiment (BJ-hTERT validation) |
| 9-24 | Phase A data collection (experiments Aim A.1, A.2, A.3) |

Setup occupies Months 1-2 of Phase A budget. Experiments begin Month 3. Data analysis through Month 6.

## 🔬 Technical specifications

### Imaging parameters

| Parameter | Value |
| --- | --- |
| Resolution | 2448×2048 px (FLIR BFS-U3-63S4M) |
| Bit depth | 12-bit → 16-bit stored |
| Frame rate | 74 FPS max (not limiting for time-lapse) |
| Exposure | 50-500 ms (brightfield), 500-3000 ms (fluorescence) |
| Z-stack | 5-11 slices × 2 μm spacing (optional for 3D imaging) |
| Channels | Brightfield + FITC (GT335/polyE) + TRITC (Ninein) |
| Time-lapse interval | 30 min (day/night) |
| Data per session | ~5 GB/day (if 10 positions × 3 channels × 48 timepoints) |
| Total data 6 mo | ~900 GB (easily fits NAS 4TB) |

### Stage automation

- XY range: Ensure coverage of 35mm dish (~40mm × 40mm) or 6-well plate

- XY accuracy: ±5 μm positioning (achievable with Arduino CNC-style steppers + linear rails; Prior stages are ±0.1 μm but more expensive)

- Z range: 10mm travel, ±0.5 μm accuracy (adequate for autofocus)

- Autofocus: Software-based, contrast detection on DAPI/Hoechst channel OR hardware laser-based (more expensive, Перспектива в upgrade)

### Environmental control

- Temperature: 37°C ±0.3°C (Peltier + PID controller)

- CO₂: 5% ±0.2% (mass flow controller + CO₂ tank; one tank lasts ~2 weeks)

- Humidity: 95% (passive evaporation from humidifier insert)

## 🌐 Remote monitoring setup

- VPN (WireGuard) into lab network from Jaba’s home

- Micro-Manager live display via VNC/RustDesk

- Camera preview stream via MJPEG over HTTP (for quick peek)

- SMS alerts via Twilio API ($0.01/msg, ~$1/month) when:

  - Temperature deviates >0.5°C

  - CO₂ deviates >0.5%

  - Camera disconnects

  - UPS engaged (power failure)

- Daily email summary — automated Python script sends overview image + environmental log at 8am

## 🧰 Список для закупки (ordering list)

### Arduino / electronics (AliExpress / RobotDyn / Amazon — $500 total)

- Arduino Mega 2560 R3 ($25)

- CNC Shield V3 ($10)

- 3× NEMA-17 stepper motors (X, Y, Z) ($40)

- 3× A4988 stepper drivers (already on shield)

- Linear rails 400mm × 2 + 200mm × 2 (X, Y) with bearings + belt drive ($120)

- 12V 10A power supply ($30)

- Endstops × 6 ($10)

- Wires + breadboard + misc ($30)

- DHT22 temperature/humidity sensor ($10)

- CO₂ solenoid valve 12V ($40)

- MH-Z19B CO₂ sensor ($25)

- DS18B20 waterproof temp probe ($10)

- MOSFET modules × 4 ($20)

- Peltier heater module + heatsink ($40)

- 3D-printed parts (stage adapter, focus adapter, chamber mount) — print yourself or Shapeways ($90)

### Camera + optics ($1,700 total)

- FLIR Blackfly S USB3 BFS-U3-63S4M-C monochrome scientific camera ($1,200, new)

- C-mount to Zeiss IM 35 adapter ($80)

- ThorLabs M470L4 blue LED (470 nm, for FITC excitation) ($180)

- ThorLabs M565L3 green LED (565 nm, for TRITC excitation) ($180)

- ThorLabs LEDD1B LED driver ($60)

### Live cell environmental ($1,050 total)

- Acrylic chamber (custom-cut 4mm thick, ~200mm × 200mm × 60mm internal, $120)

- Glass cover slip for bottom ($20)

- CO₂ tank 9kg (refillable, $60 Tbilisi medical supply)

- CO₂ regulator + flow meter ($80)

- Peltier element TEC1-12706 + heatsink + fan ($25)

- PID controller (Inkbird ITC-100 or equivalent) ($35)

- Humidifier insert (plastic tray with foam) ($10)

- OKolab compact CO₂ incubator mini chamber (used market, optional upgrade, $700)

### Computer + storage ($1,000 total)

- Refurbished Dell OptiPlex (i5-8500, 16GB, 512GB SSD, Win10 Pro) ($300)

- Synology DS220+ NAS ($230) + 2× 4TB Seagate IronWolf ($180) = $590

- APC Back-UPS Pro 1500VA ($200)

- External monitor 24” (existing or $150)

Total Variant A: $4,250 + shipping/VAT (~$200) = $4,450

## 📝 LOI v24 budget amendment

Текущая Phase A budget: $80,000

Предлагаемое изменение:

- Уменьшить “General Consumables” с $12,000 на $9,500 (-$2,500)

- Увеличить “Microscope Upgrade” с $2,000 на $4,500 (+$2,500)

- Переименовать line в “Microscope Automation & Upgrade”

- Добавить footnote: “Includes motorized XY + Z stage for round-the-clock imaging, FLIR digital camera, LED fluorescence illumination, environmental chamber (37°C + 5% CO₂), NAS backup, and UPS. Full specification in supplementary materials.”

Net effect: $80,000 total unchanged. Более точная allocation для automation.

## ⚠️ Key risks & mitigations

- DIY stage calibration drift → use linear rails с encoders (add ~$100); alignment every 2 weeks

- CO₂ tank run-out mid-experiment → install low-pressure sensor + SMS alert; keep 2 tanks rotating

- Power outage > UPS runtime → UPS gives 2h; manually restart if needed; alert sent immediately

- Contamination (long-term cultures) → weekly media change, sterile technique, monthly bleach decontamination of chamber

- Camera failure → budget $500 contingency for replacement (not in main budget; covered by $300 setup contingency + technician hours)

- Software crash → Micro-Manager + watchdog script: if no new frame in 2h, restart acquisition automatically

- Remote access compromised → WireGuard with ed25519 keys; no public SSH; lab IP whitelist

## 🤖 Claude Code /overnight — AI-operated microscopy

### Концепция

Microscope управляется Claude Code (Anthropic’s official CLI для AI-assisted workflows) в специальном /overnight режиме — автономный цикл наблюдения, принятия решений и корректировки эксперимента ночью, когда человек не в лаборатории (20:00–08:00).

В /overnight режиме agent получает авторизацию на routine decisions (выбор поле зрения, autofocus, channel switching, ROI refinement) без ежеминутного одобрения человека. Strategic decisions (изменение протокола, остановка эксперимента) остаются у человека.

### Архитектура

┌──────────────────────────────────────────────────────────────────┐
│              HUMAN (Jaba, 20:00–08:00 off-duty)                   │
│                      ↕ SMS/email alerts only                       │
└──────────────────────────────────────────────────────────────────┘
                                  ↕
┌──────────────────────────────────────────────────────────────────┐
│              CLAUDE CODE — /overnight agent                        │
│  • Reads experiment protocol (PROMPT.md, см. ниже)                 │
│  • Monitors image stream в real-time                               │
│  • Calls tool functions: move_stage(), autofocus(),                │
│     capture_multichannel(), switch_led(), etc.                     │
│  • Decision loop: каждые 30 мин inspect last batch,                │
│     выбрать next action (stay/move/rotate/alarm)                   │
│  • Journaled decisions с rationale                                 │
│  • On blocker: pauses + SMS to experimenter                        │
└──────────────────────────────────────────────────────────────────┘
                                  ↕ Python tool functions (PyMMCore-Plus)
┌──────────────────────────────────────────────────────────────────┐
│     MICRO-MANAGER 2.0 core + pymmcore-plus + Arduino firmware      │
└──────────────────────────────────────────────────────────────────┘
                                  ↕
┌──────────────────────────────────────────────────────────────────┐
│      Zeiss IM 35 + motorized XY/Z + camera + chamber + LEDs        │
└──────────────────────────────────────────────────────────────────┘

### Tool functions для Claude Code

# микроскоп-управление

move_stage_xy(x_mm, y_mm); move_stage_relative(dx, dy)
autofocus(mode='contrast'|'laplacian')
capture_single(channel, exposure_ms); capture_zstack(ch, z_range, step)
capture_multichannel(channels)
switch_led(channel, intensity_pct)

# image analysis

detect_cells(image); detect_centrioles(image)
measure_polyglu_intensity(cell_roi, channel)
segment_nuclei(image); count_dead_cells(image)

# environmental

get_temperature_c(); get_co2_percent(); get_humidity_percent()
set_target_temp(c); open_co2_valve(duration_s)

# alerting / journaling

send_sms_to_human(message, priority='info'|'warn'|'crit')
log_decision(action, rationale, observed)
save_checkpoint(state); halt_experiment(reason)  # requires human unlock

### Cost

Claude Code subscription: $20/мес Pro tier × 6 мес = $120 total — negligible vs $4,500 hardware.

## 📝 PROMPT-based experimentor signaling

### Концепция

Experimentator (Jaba) пишет PROMPT на естественном языке описывающий цели и задачи эксперимента. Claude Code интерпретирует этот prompt и автоматически:

- Мониторит данные в реальном времени (image stream, environmental log)

- Проверяет цели каждые 30 минут

- Отправляет сигналы человеку когда происходит что-то важное: достижение цели, anomaly, риск для эксперимента, необходимость intervention

Преимущество: Jaba не нужно программировать условия срабатывания. Он описывает цели словами — Claude сам формулирует критерии и мониторит.

### Формат PROMPT файла

~/Documents/Experiments/CDATA_PhaseA_Aim1/PROMPT.md:

# Aim A.1 — α vs β discrimination

## Цель

Disentangle division-dependent (α) from time-dependent (β) polyGlu accumulation
in BJ-hTERT fibroblasts by comparing proliferating vs contact-inhibited conditions.

## Условия

- **Proliferating wells** (A1-A6): sub-confluent, passaged every 3 days
- **Contact-inhibited wells** (B1-B6): seeded confluent, media changed 2× per week
- **Time points for imaging**: days 0, 7, 14, 21
- **Channels**: Brightfield, FITC (GT335 polyGlu), TRITC (Ninein mother centriole)

## Что измеряем

- polyGlu intensity at mother centriole (FITC co-localized with TRITC)
- cell density (brightfield segmentation)
- morphology changes (senescence markers: flattened cells, vacuoles)

## Сигналы, которые я хочу получить

### Позитивные (progress toward goal)

- **"Signal detected"**: если polyGlu intensity в proliferating wells вырос на >30% vs day 0 baseline
- **"α > 0 confirmed"**: если при day 14 in proliferating wells средний intensity significantly higher (p<0.05, t-test) vs contact-inhibited wells
- **"Parrinello signal"**: если 20% O₂ wells показывают >40% difference vs 3% O₂ wells (for O₂ modulation Aim A.2)

### Предупреждения (possible trouble)

- **"Contamination risk"**: если cell density drops >20% unexpectedly в 2+ wells
- **"Focus drift"**: если autofocus fails в 3+ positions consecutively
- **"Channel saturation"**: если FITC mean intensity > 60,000 counts (max 65,535)
- **"Morphology anomaly"**: если >15% cells show senescent markers earlier than expected

### Критические (требуют немедленного вмешательства)

- **"Experiment compromised"**: если cell death >40% in any well → stop imaging, SMS Jaba, save last data
- **"Environment failure"**: если temp <35°C, CO₂ <4% или >6%, humidity <80% — pause, SMS Jaba
- **"Hardware fault"**: если camera disconnects, stage jams, UPS engaged — halt, SMS Jaba

## Что Claude может делать автономно

- Выбирать best focal plane для imaging
- Adjust LED intensity within ±20% if signal too dim/bright
- Re-image if blur detected
- Skip dead cells в analysis
- Increase imaging frequency на interesting positions (if Claude detects cell division event)
- Request Z-stack если 2D image недостаточно

## Что Claude НЕ делает без моего разрешения

- Менять protocol (пример: не добавляет лишние channels)
- Остановить experiment раньше scheduled end
- Открыть chamber (sterility)
- Изменить cells физически (no chemical addition)

## Формат отчётов

- **Daily 8 AM email**: summary of last 24 hrs + 4 sample images
- **Mid-day SMS (14:00)**: short status (OK / warning / critical)
- **Cycle log (every 30 min)**: в JOURNAL.md, no notification to human unless signaled above
- **End-of-experiment report**: day 21, full analysis, recommended next steps

## Контекст для Claude (для понимания важности сигналов)

Данный эксперимент — Aim A.1 of Impetus Phase A ($80k grant). Результат определяет
go/no-go decision для Phase B ($120k). CDATA hypothesis предсказывает α > 0 значимо.
Null result (α ≈ 0 и β ≈ 0) = хonest falsification, publish negative result.

### Как Claude использует PROMPT

- Парсинг целей: извлекает из текста конкретные метрики (polyGlu intensity, cell density, effect sizes) и thresholds (>30%, >40%, p<0.05)

- Постоянная проверка: каждые 30 мин на основе последних imaging data проверяет — сработало ли какое-то условие из “Сигналы”

- Генерация сигнала:

  - “Signal detected” → send_sms с кратким summary и attached preview image

  - “Contamination risk” → send_sms + предлагает actions (“recommend: check wells A3, A5”)

  - “Experiment compromised” → halt + SMS + save checkpoint

- Log of reasoning: каждый decision журналируется в JOURNAL.md с quoting the relevant prompt line и показателями

### Пример SMS сигналов

Scenario 1 — Positive progress:

[CDATA-A1] 14:30 | INFO
α > 0 confirmed: proliferating wells show 2.3x higher polyGlu vs contact-inhibited
at day 14 (p=0.003, n=6 wells each). Keeping protocol. Full data: dashboard/cycle_156

Scenario 2 — Warning:

[CDATA-A1] 03:15 | WARN
Cell density dropped 28% in wells A4, A5 between 02:00 and 03:00 imaging.
Possible contamination. Recommend visual check at 8 AM.
Wells A1-A3, A6 normal. Continuing schedule, flagged for review.

Scenario 3 — Critical:

[CDATA-A1] 05:42 | CRIT
CO₂ dropped to 3.1% (target 5.0%). Opened valve 45s × 3, no recovery.
Tank may be empty. HALT imaging. Cells survive ~4 hrs at low CO₂.
Action needed before 10 AM. Checkpoint saved.

### Как adaptировать PROMPT для разных aim’ов

Aim A.2 (20% vs 3% O₂): добавить signals для differential wells; modify critical thresholds для O₂-sensitive markers.

Aim A.3 (CCP1/TTLL6-OE causal): добавить signals for population doubling rates, transduction efficiency (GFP% cells), senescence onset timing.

Phase B (HSC mouse): полностью другой PROMPT.md — focus на chimerism, flow cytometry results, не microscopy.

### Template для experimentor

# [Aim Name]

## Цель

[один параграф о scientific goal]

## Условия

[список conditions: wells, channels, time points]

## Что измеряем

[metrics to quantify]

## Сигналы

### Позитивные (progress)

- [condition] → signal "[name]": [action]

### Предупреждения

- [condition] → signal "[name]": [action]

### Критические

- [condition] → signal "[name]": HALT + SMS Jaba

## Автономные действия Claude

[что можно без authorization]

## НЕ делать без разрешения

[что требует human approval]

## Отчёты

- [частота: daily/mid-day/end-of-experiment]
- [формат: email/SMS/dashboard]

## Контекст

[почему это важно, stakes, acceptable null result]

### Преимущества prompt-driven approach

- Zero programming: Jaba описывает словами, не пишет Python

- Domain-expert friendly: biologist формулирует условия в биологических терминах, Claude переводит в thresholds

- Easy iteration: менять цели эксперимента = редактировать PROMPT.md, без code changes

- Explanatory: Claude’s decisions всегда ссылаются на конкретные lines PROMPT.md — transparency

- Reusable: PROMPT.md template для каждого Aim, можно делиться с коллегами

- Fail-safe: если PROMPT не покрывает сценарий — Claude не действует autонomously, pause и запрашивает human

### Validation перед submission

Перед тем как Claude запускает /overnight, Jaba делает:

- Dry run на синтетических данных — проверить что сигналы срабатывают корректно

- Day-1 validation: первые 24 часа Claude в /overnight режиме, но Jaba проверяет logs каждый час вручную

- Progressive автономность: после 48 часов без issues — full autonomy; если issue — уменьшить autonomy

### Для Impetus LOI v24 — Methodological Innovation

Этот prompt-driven signaling — дополнительный scientific selling point:

“Our automated microscopy platform pairs low-cost retrofit ($4,500) with prompt-driven AI supervision (Claude Code /overnight). The experimenter articulates goals и thresholds в natural-language PROMPT.md; Claude continuously monitors data and signals the human only when strategically important events occur (positive progress, warnings, critical failures). This eliminates the ‘experimenter wake-up cost’ traditional time-lapse requires and democratizes AI-assisted experimental science for resource-limited labs. All PROMPT templates, tool functions, and decision logs released under MIT license concurrent with Phase A preprint.”

## 🚀 Follow-up upgrades (post-Phase A, if Go decision)

If Phase A passes и funding extends (Phase B or follow-on grant):

- Hamamatsu ORCA-Fusion BT upgrade (+$4k used) → higher sensitivity for low-expression fluorescence

- Perfect Focus System (Nikon PFS) retrofit (+$3k used) → hardware-based autofocus, no software drift

- Motorized filter turret (+$1k) → automated channel switching, faster imaging

- Liquid handling robot (OpenTrons OT-2 $2.5k new) → automated media change, ideal for 3-4 week contact-inhibition protocols

Budget for upgrades: place into Phase B or post-grant $10-15k capital line.

## ✅ Decision for LOI v24 submission (2026-04-24)

Recommendation: Вариант A ($4,500 total equipment line) — максимально cost-effective, polностью покрывает Phase A experimental requirements (automated day-night imaging, environmental control, remote monitoring, data backup). Upgradeable to Phase B levels later.

Action:

- Amend Phase A budget: “Microscope Upgrade” $2k → “Microscope Automation & Upgrade” $4.5k (reallocate $2.5k from Consumables)

- Add this document as supplementary material

- Include in Cover Letter: “A custom-engineered automated time-lapse imaging platform (see Supplementary: Automated Microscopy Setup) will enable round-the-clock tracking of centriolar polyGlu dynamics in live fibroblasts, maximizing temporal resolution within budget constraints.”

Engineered 2026-04-21 for Impetus LOI v24 submission 2026-04-25.